1/23/2004

MissPiggy Thermometers and Heaters

The colored sensors and heaters are normally connected to the counting room

| 1 | A | I- | Diffuser Temp. Sensor GaAlAs(1) | TG-120P | 8792 Not Calib. |
| --- | --- | --- | --- | --- | --- |
|  | B | V+ | GaAlAs(1) |  |  |
|  | C | V- | GaAlAs(1) |  |  |
|  | D | V- | Diffuser Temp. Sensor GaAlAs(2) | TG-120P | 9279 Calib. |
|  | E | V+ | GaAlAs(2) |  |  |
|  | F | I+ | GaAlAs(2) |  |  |
|  | H | I+ | GaAlAs(1) |  |  |
|  | J | I- | GaAlAs(2) |  |  |
|  | K |  | Diffuser Heater | Manganin wire | 43 Ω |
|  | L |  |  |  |  |

| 2 | A | I- | Phase Separator Si(470) | DT-470-BO-13 |  |
| --- | --- | --- | --- | --- | --- |
|  | B | V+ | Phase Separator Si(470) |  |  |
|  | C | V- | Phase Separator Si(470) |  |  |
|  | D | I- V- | Sample Tube Inner Shield | DT-471-SD |  |
|  | E | I+ V+ | Sample Tube Inner Shield |  |  |
|  | F | I+ V+ | Phase Sept. Outer Shield | DT-471_BO |  |
|  | H | I+ | Phase Separator Si(470) |  |  |
|  | J | I- V- | Phase Sept. Outer Shield |  |  |
|  | K |  | Phase Sept. Heater | 2 Watlow Kapton 1 by 3 inch | 26 Ω |
|  | L |  |  |  |  |

| 3 | A | I- V- | Temp. Gas Entering Trans. Line | DT-471-BO |  |
| --- | --- | --- | --- | --- | --- |
|  | B | I+ V+ | T. Sample Tube Window COLD | Ptr |  |
|  | C | I- V- | T. Sample Tube Window COLD |  |  |
|  | D | I- V- | Temp. Sample Tube WARM | Ptr |  |
|  | E | I+ V+ | Temp. Sample Tube WARM |  |  |
|  | F |  | Sample Tube Heater WARM | 2 Watlow Kapton | 55 Ω |
|  | H | I+ V+ | Temp. Gas Entering Trans. Line |  |  |
|  | J |  | Sample Tube Heater WARM |  |  |
|  | K |  | Sample Tube Heater Window COLD | 2 Watlow Kapton 5 by 5 inch2 | 55 Ω |
|  | L |  |  |  |  |